Interferential Current Stimulation

Student handout — Block B  ·  Physics of Electrical Neurostimulation

1st FALAN Latin American Training Program in Neuroscience — Santiago, August 2026

NeuroEng@USACH · Leonel Medina, Rodrigo Osorio, Cristián Morales

Name: _______________________________     Partner(s): _______________________________

Volunteer initials: __________   Date: __________   Consent confirmed by: __________

Companion notebook: notebooks/02_interferential_stimulation.ipynb Safety, exclusion criteria and consent script: SAFETY.md — the same rules as Block A apply.

# By the end of this session you will be able to

- Explain how two medium-frequency (kHz) currents interfere to produce a low-frequency amplitude-modulated beat inside the tissue.

- Identify where the beat amplitude is largest for a given montage, and set up a cross (quadripolar) configuration on the forearm.

- Elicit pulsatile finger-muscle contractions with a carrier that is itself too fast for the axon to follow pulse by pulse.

- Demonstrate the block of cutaneous afferents under a kHz field, using a plastic probe on the skin between the electrodes.

- Contrast interferential stimulation with the single-pulse stimulation of Block A.

## Session map (70 min)

| Part | Time |
| --- | --- |
| Module 0 — how interferential current works (notebook Parts 1–2) | ~15 min |
| Module 1 — the field of four electrodes; where the beat is deepest (Part 3) | ~20 min |
| Module 2 — cross montage, motor threshold and pulsatile contraction | ~20 min |
| Module 3 — cutaneous afferent block with the plastic probe | ~10 min |
| Wrap-up and discussion | ~5 min |

## What you need

A two-channel interferential stimulator (CH1 and CH2 independently programmable in the kHz range) · four self-adhesive electrodes · conductive gel · alcohol wipes · a blunt plastic probe · skin marker · stopwatch · a volunteer.

# Module 0 — How interferential current works

## Background

Two sinusoidal currents at slightly different medium frequencies — here f1 = 4000 Hz on CH1 and f2 = 4002 Hz on CH2 — are injected through separate electrode pairs. Where the two fields overlap inside the tissue they add, and the sum is a ~4 kHz carrier whose amplitude swells and fades at the difference frequency Δf = |f2 − f1|. That slow envelope is the stimulus the nerve responds to, and it is generated inside the tissue rather than applied at the surface.

The axon membrane behaves as a low-pass filter with a time constant of roughly 0.1–1 ms. A 4 kHz carrier is far too fast to be followed cycle by cycle, so each individual cycle contributes almost no net depolarisation — but the envelope is slow enough to drive the membrane.

Because the carrier is medium-frequency, skin impedance is low (Z falls with frequency), so the stimulus is comfortable and reaches deeper than a low-frequency pulse of the same perceived strength.

Continuous kHz current also depolarises and then desensitises the fine cutaneous afferents directly under the electrodes — carrier-induced conduction block. That is what Module 3 tests.

## From the notebook, before you touch hardware

Run Part 2 at Δf = 100 Hz and at Δf = 2 Hz. Then run Part 3 and record the number that matters:

Modulation depth at the geometric centre of the montage: ______ %

Peak modulation depth on the plane, and roughly where it sits: ______ % at ______________

Predict before running Part 3: where will the beat be deepest? _____________________

Were you right? If not, what is wrong with the usual picture of the crossing point as the "focus"?

______________________________________________________________________________

______________________________________________________________________________

# Module 1 — The cross (quadripolar) montage

## Procedure — electrode placement

- Seat the volunteer with the forearm supported, palm up, wrist and fingers relaxed and free to move. Clean the skin with alcohol and let it dry.

- Identify the target: the muscle belly of the finger flexors on the volar (anterior) forearm, about one third of the way down from the elbow crease.

- CH1, parallel to the nerve: place its two electrodes along the proximal–distal axis, one proximal and one distal, so its current runs along the nerve/muscle direction.

- CH2, perpendicular to the nerve: place its two electrodes on the medial and lateral sides of the same region, so its current crosses CH1 at roughly 90°.

- The four electrodes form a square or diamond with the two current paths crossing near the centre, over the finger flexors.

- Keep at least 3–4 cm between the two electrodes within a channel, and press each electrode flat with no air gaps. Mark the centre of the cross lightly with the skin marker.

- Have a facilitator confirm the montage before switching on.

## Task 1 — Set the stimulator

Program the following and record what your device actually displays.

| Parameter | Set value | Displayed / notes |
| --- | --- | --- |
| CH1 carrier frequency | 4000 Hz |  |
| CH2 carrier frequency | 4002 Hz |  |
| Beat (difference) frequency | = ______ Hz |  |
| Waveform | sinusoidal, continuous |  |
| CH1 amplitude at motor threshold | ______ mA |  |
| CH2 amplitude at motor threshold | ______ mA |  |
| Electrode size / type |  |  |

Both channels must be raised together. The beat exists only where the two fields overlap with comparable amplitude. If one channel is much stronger than the other, the montage degenerates into ordinary bipolar stimulation from the stronger pair and there is no interference to speak of.

# Module 2 — Motor threshold and pulsatile contraction

Goal: bring the volunteer from no sensation, through tingling, to a visible pulsatile finger-muscle contraction driven by the 2 Hz beat.

## Procedure

- Start both channels at 0 mA. Ramp CH1 and CH2 together in ~1 mA steps, pausing 2–3 s at each step and asking the volunteer what they feel.

- Record three landmarks: (a) sensory threshold — first tingling; (b) motor threshold — first visible twitch of the fingers; (c) comfortable pulsatile contraction.

- Count the contractions over 10 s and divide by 10. Compare with Δf = f2 − f1.

- Note which fingers move. The finger flexors are innervated by the median nerve, and the ulnar nerve for the 4th–5th digits, so the pattern tells you where the beat actually landed.

- Detune the beat: set CH2 to 4050 Hz (Δf = 50 Hz) and try to reach motor threshold again. Record what happens to the character of the contraction.

- Finally set CH2 to 4000 Hz (Δf = 0). Record whether any contraction persists.

| Condition | CH1 / CH2 (Hz) | Δf (Hz) | Amplitude (mA) | What you observed |
| --- | --- | --- | --- | --- |
| Sensory threshold | 4000 / 4002 | 2 |  |  |
| Motor threshold | 4000 / 4002 | 2 |  |  |
| Comfortable pulsatile contraction | 4000 / 4002 | 2 |  |  |
| Detuned beat | 4000 / 4050 | 50 |  |  |
| No beat | 4000 / 4000 | 0 |  |  |
| Single channel only (CH1 on, CH2 off) | 4000 / — | — |  |  |

Measured contractions per second at Δf = 2 Hz: ______

## Question

The carrier delivers 4000 cycles per second, yet the muscle contracts only about twice per second. Explain this in terms of membrane time constant and low-pass filtering.

______________________________________________________________________________

______________________________________________________________________________

# Module 3 — Blocking the cutaneous afferents

Goal: while the muscle is still contracting to the beat, show that light-touch sensation on the skin inside the cross is reduced compared with skin outside it.

## Procedure

- Control, before switching on. With the stimulator at 0 mA, touch the volunteer with the blunt plastic probe at four points and take a 0–10 rating at each.

- The volunteer keeps their eyes closed for all probe testing, and the tester varies the order of sites so ratings are not anticipated.

- Ramp both channels back up to the comfortable pulsatile contraction from Module 2 and hold it steady for about 60 s.

- With stimulation running, repeat the probe test at A, B, C and D in a different order.

- Take the current to 0 and retest all four sites immediately, then again after 2 min, to see how quickly sensation recovers.

- Swap roles and repeat with a second volunteer if time allows.

| Probe site | Before (0–10) | During (0–10) | Just after (0–10) | +2 min (0–10) |
| --- | --- | --- | --- | --- |
| A — centre of the cross (between electrodes) |  |  |  |  |
| B — 2 cm outside the montage, same side |  |  |  |  |
| C — dorsal forearm, same arm |  |  |  |  |
| D — opposite forearm (control) |  |  |  |  |

Was the muscle still contracting while site A felt blunted?   Yes  /  No

Block index = (rating at B during) − (rating at A during) = ______

## Question

Site A is closer to the electrodes than site B, yet the motor axons under A are being driven vigorously. Why can the same field block small superficial sensory fibres while continuing to drive deeper large motor axons?

______________________________________________________________________________

______________________________________________________________________________

# Comparison with Block A

|  | Single-pulse stimulation (Block A) | Interferential stimulation |
| --- | --- | --- |
| Frequency content |  |  |
| Current at motor threshold (mA) |  |  |
| Reported comfort (0–10) |  |  |
| Contraction rate (per second) |  |  |
| Touch sensation between electrodes |  |  |

## Discussion

- Did the contraction appear at the crossing point, where the notebook says the beat is weakest? If it did, what does that tell you about the real current paths — think about the anisotropy of muscle, the symmetry of your montage, and the depth of the nerve relative to the plane the notebook plotted.

- Rotating one channel moves where the crossing point falls. If you moved CH2 by 2 cm proximally, which fingers would you expect to respond, and why?

- How would you separate a genuine kHz block of the afferents from simple habituation, or from distraction by the ongoing contraction? Design one extra control you could run in five minutes.

# Wrap-up questions

1. Write the beat frequency in terms of f1 and f2, and state what changes physically inside the tissue when Δf goes from 2 Hz to 50 Hz.

______________________________________________________________________________

______________________________________________________________________________

2. Why does a 4 kHz carrier feel more comfortable than a 100 Hz pulse train at the same perceived strength?

______________________________________________________________________________

______________________________________________________________________________

3. The crossing point of the two channels is often described as the place where stimulation is "focused". Based on what you observed and on the field model, how would you correct that statement?

______________________________________________________________________________

______________________________________________________________________________

# Quick reference

- Beat (difference) frequency: Δf = f2 − f1 → 4002 − 4000 = 2 Hz. - Effective carrier: (f1 + f2)/2 ≈ 4001 Hz; envelope amplitude ∝ 2·A·cos(π·Δf·t). - Modulation depth m = ( ¦a + b¦ − ¦a − b¦ ) / ( ¦a + b¦ + ¦a − b¦ ), where a and b are the two channels' current-density vectors. m = 1 where they are collinear and equal; m = 0 where they are perpendicular and equal — which is exactly the geometric centre of a symmetric cross montage. The deep modulation sits in four off-centre lobes. - Membrane low-pass filtering: τm ≈ 0.1–1 ms. The 125 µs half-cycle at 4 kHz is comparable to the chronaxie of Aα motor fibres (~50–100 µs) but delivers little net depolarisation per cycle. - Skin impedance falls with frequency, so kHz carriers are more comfortable and penetrate deeper than low-frequency pulses. - kHz conduction block: sustained medium-frequency current desensitises small superficial cutaneous afferents while deeper large motor axons still follow the envelope. - Today's setting: CH1 4000 Hz, CH2 4002 Hz, sinusoidal continuous, cross (quadripolar) montage over the volar forearm finger flexors.